## Sheet1
CCB GPIO-104 INPUT & OUTPUT SUMMARY
CCB GPIO-104 ANALOG INPUTS
INTERNAL | RANGE OF ALLOWABLE VALUES | [5 empty] | ANALOG INPUT PARAMETERS
POINT BEING
MONITORED | MINIMUM | NOMINAL | MAXIMUM | INPUT | FULL SCALE | FULL SCALE | MINIMUM | MINIMUM | NOMINAL | NOMINAL | MAXIMUM | MAXIMUM
| ( VOLTS ) | ( VOLTS ) | ( VOLTS ) | CHANNEL | ( VOLTS ) | ( A/D COUNTS ) | ( VOLTS ) | ( A/D COUNTS ) | ( VOLTS ) | ( A/D COUNTS ) | ( VOLTS ) | ( A/D COUNTS )
+5V (DIG) (SEE NOTE 2) | 4.75 | 5 | 5.25 | Ain 7 | 5 | 4095 | 2.375 | 1945.125 | 2.5 | 2047 | 2.625 | 2149
+8V (ANA) (SEE NOTE 2) | 7.6 | 8 | 8.4 | Ain 6 | 5 | 4095 | 2.375 | 1945.125 | 2.5 | 2047 | 2.625 | 2149
+12V (FAN) (SEE NOTE 2) | 11.4 | 12 | 12.6 | Ain 5 | 5 | 4095 | 2.375 | 1945.125 | 2.5 | 2047 | 2.625 | 2149
SM HEARTBEAT | 1.32 | 1.65 | 1.98 | Ain 4 | 5 | 4095 | 2 | 1638 | 2.5 | 2047 | 3 | 2457
+5.0VA | 4.75 | 5 | 5.25 | Ain 3 | 5 | 4095 | 2.375 | 1945.125 | 2.5 | 2047 | 2.625 | 2149
+3.3VD | 3.135 | 3.3 | 3.465 | Ain 2 | 5 | 4095 | 2.375 | 1945.125 | 2.5 | 2047 | 2.625 | 2149
+2.5VD | 2.375 | 2.5 | 2.625 | Ain 1 | 5 | 4095 | 2.375 | 1945.125 | 2.5 | 2047 | 2.625 | 2149
+1.2VD | 1.14 | 1.2 | 1.26 | Ain 0 | 5 | 4095 | 2.375 | 1945.125 | 2.5 | 2047 | 2.625 | 2149
NOTE 1) All A/D input channels are scaled for +5.000Vdc FULL SCALE. All input measurands are scaled to present a NOMINAL voltage of +2.500Vdc ( 1/2 FULL SCALE ) under normal operating conditions.
NOTE 2) These monitored points are NOT multiplexed. ( All other monitored points ARE multiplexed via the MON SEL bus and report back as MONITOR PORTS 0, 1, 2, 3, and 7 with PORTS 4, 5, and 6 left available for future enhancements )
MONITOR PORT MAP: | PORT 0 = Backplane Daughter Card Slot 0 ( leftmost )
PORT 1 = Backplane Daughter Card Slot 1
PORT 2 = Backplane Daughter Card Slot 2
PORT 3 = Backplane Daughter Card Slot 3 ( rightmost )
PORT 4 = ( SPARE )
PORT 5 = ( SPARE )
( Digital portion ONLY ) | PORT 6 = IP Addess ID ( Used to read switches installed on CONSOLE / TEST PORT cable to establish Ethernet IP address during installation ) ( "/CFG ERROR" = ID0, "/CFG DONE" = ID1 )
PORT 7 = Master Card ( Backplane )
CCB GPIO-104 DIGITAL I/O
82C55 PORT / BIT | MODE | DIRECTION | SIGNAL NAME |  |  |  |  | DESCRIPTION / NOTES
A-0 | 0 ( SIMPLE I/O ) | OUTPUT | SAMPLE |  |  |  |  | This signal is set HIGH to indicate that the CCB is in its SAMPLE MODE
A-1 | 0 ( SIMPLE I/O ) | OUTPUT | FPGA 4 DONE |  |  |  |  | This signal is set HIGH to indicate that FPGA 4 has completed loading its configuration bitstream
A-2 | 0 ( SIMPLE I/O ) | OUTPUT | FPGA 3 DONE |  |  |  |  | This signal is set HIGH to indicate that FPGA 3 has completed loading its configuration bitstream
A-3 | 0 ( SIMPLE I/O ) | OUTPUT | FPGA 2 DONE |  |  |  |  | This signal is set HIGH to indicate that FPGA 2 has completed loading its configuration bitstream
A-4 | 0 ( SIMPLE I/O ) | OUTPUT | FPGA 1 DONE |  |  |  |  | This signal is set HIGH to indicate that FPGA 1 has completed loading its configuration bitstream
A-5 | 0 ( SIMPLE I/O ) | OUTPUT | FPGA 0 DONE |  |  |  |  | This signal is set HIGH to indicate that FPGA 0 has completed loading its configuration bitstream
A-6 | 0 ( SIMPLE I/O ) | OUTPUT | SYSTEM OK |  |  |  |  | This signal is set HIGH to indicate that all FPGA's present have healthy heartbeats ( See "CCB GPIO-104 ANALOG INPUTS" section above )
A-7 | 0 ( SIMPLE I/O ) | OUTPUT | PWR OK |  |  |  |  | This signal is set HIGH to indicate that all power supplies are operating within acceptable limits ( See "CCB GPIO-104 ANALOG INPUTS" section above )
B-0 | 0 ( SIMPLE I/O ) | OUTPUT | FPGA 4 ERROR |  |  |  |  | This signal is set HIGH to indicate that FPGA 4 has encountered an error condition while loading its configuration bitstream
B-1 | 0 ( SIMPLE I/O ) | OUTPUT | FPGA 3 ERROR |  |  |  |  | This signal is set HIGH to indicate that FPGA 3 has encountered an error condition while loading its configuration bitstream
B-2 | 0 ( SIMPLE I/O ) | OUTPUT | FPGA 2 ERROR |  |  |  |  | This signal is set HIGH to indicate that FPGA 2 has encountered an error condition while loading its configuration bitstream
B-3 | 0 ( SIMPLE I/O ) | OUTPUT | FPGA 1 ERROR |  |  |  |  | This signal is set HIGH to indicate that FPGA 1 has encountered an error condition while loading its configuration bitstream
B-4 | 0 ( SIMPLE I/O ) | OUTPUT | FPGA 0 ERROR |  |  |  |  | This signal is set HIGH to indicate that FPGA 0 has encountered an error condition while loading its configuration bitstream
B-5 | 0 ( SIMPLE I/O ) | OUTPUT | CALIB |  |  |  |  | This signal is set HIGH to indicate that the CCB is in its ( fixed ) CALIBRATION MODE
B-6 | 0 ( SIMPLE I/O ) | OUTPUT | DUMP |  |  |  |  | This signal is set HIGH to indicate that the CCB is in its DUMP MODE
B-7 | 0 ( SIMPLE I/O ) | OUTPUT | TEST |  |  |  |  | This signal is set HIGH to indicate that the CCB is in its TEST MODE
C-0 | 0 ( SIMPLE I/O ) | OUTPUT | MON SEL 0 |  |  |  |  | Least significant bit of the CCB's internal monitor port address bus
C-1 | 0 ( SIMPLE I/O ) | OUTPUT | MON SEL 1 |  |  |  |  | Next significant bit of the CCB's internal monitor port address bus
C-2 | 0 ( SIMPLE I/O ) | OUTPUT | MON SEL 2 |  |  |  |  | Most significant bit of the CCB's internal monitor port address bus
C-3 | 0 ( SIMPLE I/O ) | OUTPUT | /FORCE CFG |  |  |  |  | Pulse this instrument-wide signal LOW briefly to force all FPGA's to be re-configured without powering down the CCB
C-4 | 0 ( SIMPLE I/O ) | INPUT | SM /CFG ERROR |  |  |  |  | A LOW level on this multiplexed input indicates that an error was encountered while configuring the FPGA addressed by the MON SEL bus
C-5 | 0 ( SIMPLE I/O ) | INPUT | SM /CFG DONE |  |  |  |  | A LOW level on this multiplexed input indicates that the FPGA addressed by the MON SEL bus has completed loading its configuration bitstream
C-6 | 0 ( SIMPLE I/O ) | INPUT | ALL CFG DONE |  |  |  |  | A HIGH level on this open-drain, or-tied input indicates that ALL FPGA's have completed loading their configuration bitstreams
C-7 | 0 ( SIMPLE I/O ) | INPUT | HIGH TEMP |  |  |  |  | A HIGH level on this input indicates that the CCB's enclosure has reached an operating temperature equal to or greater than 50 degrees C
18-JUL-05
-RLM-