A10308D001

GBT Prime Focus 2 (0.910-1.230 GHz) RECEIVER

MONITOR/CONTROL DOCUMENT

April 1, 2002

Page 1 of 11

Rev. A  July 29, 2002  --  Several corrections from M. Mello

Rev. B  September, 2002 – Corrections from G. Anderson during troubleshooting

1)		System Description

Drawing number D10308K001 shows a block diagram of the Prime Focus 2 receiver.

In its natural mode of operation, this unit receives two linear, orthogonally polarized inputs from 0.910-1.230 GHz.  Dual circular polarization may be synthesized from the two linearly polarized inputs by switching into the system a 90 degree hybrid, attenuators, and phase shifters.

The two RF signals are amplified by cryogenic amplifiers, filtered, amplified further and routed through transfer switches which are used to select linear or circular polarization.  The signals are then mixed to a higher frequency, filtered, and amplified.  The signals (either linear or circularly polarized) are then sent to a filter bank so that various narrow or wide band filters may be selected.  Finally, variable attenuators are used to set the receiver output level.

Data may be output from the receiver in three ways:

- Coaxial output to the IF rack in the receiver room

- Digitized total power sample via the MCB

- Digitized total power sample via fiber from the Prime Focus receiver to the DCR.

All monitor and control functions are accomplished through a single VLBA MCB Standard Interface Board (SIB).  All relative addresses and their functions are listed in Table I.  The Prime Focus MCB is addressed at 41 H.

Note:  Universal definitions: TTL Hi, TTL High, and “1” are equivalent.  TTL Low, TTL Lo, and “0” are equivalent.

## Table I

Monitor and Control Functions

Rel. Hex		Description

Address

Digital Control/ Readback

48		Circular Polarizer attenuator and phase shifter for X polarization input

49		Circular Polarizer attenuator and phase shifter for Y polarization input

57, 58		Filter selection, X and Y Polarization

4C		Cryogenics Control

4D		Amplifier power control, Linear/Circular Polarization select

52, 53		Receiver Output Attenuators, X and Y polarization.

54		Noise Source Control

Digital Monitor

60 		Cryogenic state, Noise Switch Control, and Polarization monitor.

61		Receiver Identification.  Described by Table XI.

2)		System Control

2.1) 	Cryogenic State Control

Three Bits allow control of the receiver’s refrigerator and heaters.  The cryogenic state may be set by writing the desired code to relative address 4Ch, bits 7, 6, and 5.  The cryogenic control code description is given in Table 2.

Table 2

Cryogenic State Control

Relative Address 4C

B7	b6	b5	Bit 	Name		Function

X	not H	C				(from D53200A004)

1	1	0	  6	OFF		No refrigerator power, heater power

1	1	1	  7	COOL		Normal cooled operation.

1	0	1	  5	HEAT		Applies power to heaters on charcoal trap.

2.2)		Calibration Signal Control

Seven bits at relative address 54 control the noise calibration signals.  They are defined as follows:

Bit 0 -	Determines whether control for the YR Polarization channel shall come from the MCB or from an external source.  Normal operation for the switch during observation shall come from  an external source.	(0 = External, 1 = MCB)

( YRMCB/External Select from B35201S009, LO Detector/Cal Control Card Schematic)

Bit 1 -	Determines whether control for the XL Polarization channel shall come from the MCB or from an external source.  Normal operation for the switch during observation shall come from  an external source.	(0 = External, 1 = MCB)

(XLMCB/External Select:  from B35201S009, LO Detector/Cal Control Card Schematic)

Bit 2 -	Input to the Y Polarization noise cal switch (0 = Off, 1 = On)

(YRCPUNoiseSwCtrl:  from  B35201S009, LO Detector/Cal Control Card Schematic)

Bit 3 -	Input to the X Polarization noise cal switch (0 = Off, 1 = On)

(XLCPUNoiseSwCtrl:  from B35201S009, LO Detector/Cal Control Card Schematic)

Bit 4 -	Controls power to the Lo Power Noise Source.  (0 = On, 1 = Off)

(CPULoCalPwrSw:  from B35201S009, LO Detector/Cal Control Card Schematic)

Bit 5 -	Controls the switch which selects Hi or Lo cal.  Also controls power to the Hi power Noise Source, turning power on to the Noise Source when Hi Cal is selected and turning power off to the Noise Source when it is not in use.

(0 = Off (low cal), 1 = On (hi cal))

(Lo or Hi Cal Select::  from B35201S009, LO Detector/Cal Control Card Schematic)

Bit 6 -	Selects between calibration control by switches mounted inside the front end box or remote operation (0 = Remote. 1 = Local)

(Local/CPU Ctrl Select:  from B35201S009, LO Detector/Cal Control Card Schematic)

Examples:

MSB					LSB

b6	b5	b4	b3	b2	b1	b0

Code						Function description

0	0	0	0	0	0	0	00h		External Control (x and y) , LOW cal.

0	1	0	0	0	0	0	20h		External Control (x and y), HIGH cal.

0	0	0	1	0	1	0	0Ah		LOW cal selected, “X” switch is under MCB control and commanded closed, “Y” switch is under external control.

2.3)		Filter Bank Switch Control

Two bits at relative address 57h and eight bits at relative address 58h are reserved for control of the  switches used to select the desired filter bandwidth.

Table 3

Filter Selection (x polarization)

| Filter Bandwidth (x pol) | Relative Address | Bit |
| --- | --- | --- |
| External | 57 | 1 |
| 240 MHz | 57 | 0 |
| 80 MHz | 58 | 7 |
| 40 MHz | 58 | 6 |
| 20 MHz | 58 | 5 |

Table 4

Filter Selection (y polarization)

| Filter Bandwidth (y pol) | Relative Address | Bit |
| --- | --- | --- |
| External | 58 | 4 |
| 240 MHz | 58 | 3 |
| 80 MHz | 58 | 2 |
| 40 MHz | 58 | 1 |
| 20 MHz | 58 | 0 |

Each table above corresponds to a filter bank.  For each filter bank, all bits should stay high except for the bit corresponding to the desired filter, see Tables 3 and 4.  Upon power up, if no previous state was recorded in memory, the 80 MHz filter should be selected.

2.4)		Cryogenic Amplifier Control

The control word at relative address 4Dh applies or removes power to the cryogenic amplifiers.  The default setting should be “0” (amplifiers ON).

Table 5

Amplifier Bias Control

Relative Address 4Dh

Bit	Amplifier

7	X Amplifier

6	Y Amplifier

(0 = bias off,  1 = bias on)

2.5 Polarization Control

One bit is needed to control the transfer switches which allow the user to configure the receiver for either linear or circular polarization.  This one bit is located at relative address 4Dh, bit 5.  When this bit is set high the receiver will be in circular polarized mode.  When this bit is set low the receiver will be in linear polarized mode.  At  power up the receiver should default to linear mode.

2.6 Circular Polarization Settings

When circular polarization is selected, two phase shifters and two attenuators shall be initialized from a frequency dependent file in the receiver’s configuration file.  If the observation frequency is changed, the values for the two attenuators and phase shifters shall be reinitialized from the receiver configuration file.  Additionally, an operator must be able to manually enter values for each attenuator and phase shifter at the manager level.  Manually entered values will be written over by values from the configuration file if the frequency is moved.  Exception: Values shall NOT be changed during an scan.  (See the manager for Prime Focus receiver #1 for more details.)

For polarization X and Y, the attenuator and phase shifter may be set through relative addresses 48 and 49, respectively.

Table 6:  X-Polarization:

| Relative Address | Bit | Function | Sense |
| --- | --- | --- | --- |
| 48 | 13 | Pol X 123 Degree Insert | “1” = Insert Delay “0” = No Delay |
| 48 | 12 | Pol X 61.6 Degree Insert | “1” = Insert Delay “0” = No Delay |
| 48 | 11 | Pol X 30.8 Degree Insert | “1” = Insert Delay “0” = No Delay |
| 48 | 10 | Pol X 15.4 Degree Insert | “1” = Insert Delay “0” = No Delay |
| 48 | 9 | Pol X 7.7 Degree Insert | “1” = Insert Delay “0” = No Delay |
| 48 | 8 | Pol X 3.85 Degree Insert | “1” = Insert Delay “0” = No Delay |
| 48 | 7 | Pol X 1.92 Degree Insert | “1” = Insert Delay “0” = No Delay |
| 48 | 6 | Unused |  |
| 48 | 5 | Pol X 8 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 48 | 4 | Pol X 4 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 48 | 3 | Pol X 2 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 48 | 2 | Pol X 1 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 48 | 1 | Pol X 0.5 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 48 | 0 | Pol X 0.25 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |

Table 7:  Y-Polarization:

| Relative Address | Bit | Function | Sense |
| --- | --- | --- | --- |
| 49 | 13 | Pol Y 123 Degree Insert | “1” = Insert Delay “0” = No Delay |
| 49 | 12 | Pol Y 61.6 Degree Insert | “1” = Insert Delay “0” = No Delay |
| 49 | 11 | Pol Y 30.8 Degree Insert | “1” = Insert Delay “0” = No Delay |
| 49 | 10 | Pol Y 15.4 Degree Insert | “1” = Insert Delay “0” = No Delay |
| 49 | 9 | Pol Y 7.7 Degree Insert | “1” = Insert Delay “0” = No Delay |
| 49 | 8 | Pol Y 3.85 Degree Insert | “1” = Insert Delay “0” = No Delay |
| 49 | 7 | Pol Y 1.92 Degree Insert | “1” = Insert Delay “0” = No Delay |
| 49 | 6 | Unused |  |
| 49 | 5 | Pol Y 8 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 49 | 4 | Pol Y 4 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 49 | 3 | Pol Y 2 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 49 | 2 | Pol Y 1 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 49 | 1 | Pol Y 0.5 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 49 | 0 | Pol Y 0.25 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |

Multiple bits may be “ON” at any time.  The resulting attenuation or phase insertion is then the sum of the values of the attenuation or phase shifts corresponding to the individual bits.

The user should be able to enter a decimal value for attenuation or phase (i.e., 3dB attenuation or 17 degrees phase).  When entering an attenuation or phase shift value, the bits selected should produce delays whose sum most closely matches the desired value.

  - Output Attenuators

The attenuator values that follow are separate and distinct from the values given in Section 2.6 above.  These values are adjusted by the user to set the receiver output level.

Table 8:  X-Polarization:

| Relative Address | Bit | Function | Sense |
| --- | --- | --- | --- |
| 52 | 6 | Pol X 32 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 52 | 5 | Pol X 16 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 52 | 4 | Pol X 8 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 52 | 3 | Pol X 4 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 52 | 2 | Pol X 2 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 52 | 1 | Pol X 1 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 52 | 0 | Pol X 0.5 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 53 | 6 | Pol Y 32 dB Insert |  |
| 53 | 5 | Pol Y 16 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 53 | 4 | Pol Y 8 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 53 | 3 | Pol Y 4 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 53 | 2 | Pol Y 2 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 53 | 1 | Pol Y 1 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |
| 53 | 0 | Pol Y 0.5 dB Insert | “1” = Insert Attenuation “0” = No Attenuation |

Multiple bits may be “ON” at any time.  The resulting attenuation insertion is then the sum of the values of the attenuation corresponding to the individual bits.

The user should be able to enter a decimal value for attenuation (i.e., 3dB or 5.5 dB attenuation).  When entering an attenuation value, the bits selected should produce delays whose sum most closely matches the desired value.

3.0)		Monitor

3.1)		Digital Monitor

The cryogenic state and switch states are monitored at points closer to the actual hardware for fault  detection within the control circuitry.  If the monitor is not the same as the control address readback then there is a problem in the receiver circuitry.

3.1.1)	Cryogenic State Monitor

The Cryogenic state will be monitored as follows:

Table 9

# Cryogenic State Monitor

Relative Address 60

b2	b1	b0 	Name

X	H	C	(from D53200A004)

1	1	0	OFF

1	1	1	COOL

1	0	1	HEAT

0	1	1	PUMP

3.1.2) 	Switch Status Monitors

Relative address 60h also monitors the state of the noise calibration switches.

Bit 5	XL Polarization External Cal Monitor,  Monitors the state of control voltage going to the XL polarization noise cal switch.  (“1” = Cal Off, “0” = Cal On)

Bit 4	YR Polarization External Cal Monitor,  Monitors the state of control voltage going to the YR polarization noise cal switch.  (“1” = Cal Off, “0” = Cal On)

Bit 3	Polarization switch monitor.  Shows the state of the polarization switch.   A “1” indicates that the switch is in the Linear position.  A “0” indicates that the switch is in the Circular position.

3.1.2	Rack Identification

The identification of the rack can be read at relative address 61h, bits 0 through 7.  Data for this byte shall be hard-wired into connector P12 by tying all bits set at “0” to ground.  Bits set at “1” shall be left unconnected.  Bit 7 is an odd-parity bit and must be set to “1”.  The bit assignment (from high-order to low-order) is:

11000001

Bits 13 through 8, set aside for Serial Number and Modification Number, are unused.

3.2		Analog Monitor

All analog monitor data comes from the SIB as a 12 bit signed integer, left justified (A/D MSB is in integer bit 15).  After right justification and sign extension, multiply integer by 4.8828 x 10-3 V/Count to obtain volts, then multiply by the multiplier stated in Table 9.

Alarms for the analog monitor points will be specified on the installation sheet of the receiver.

Display = Mult * Counts * 4.8828e-3

Table 10:  Analog Monitor Parameters

MCB

RA				Normal	Sig.	Sample	Display	Log

(hex)	Description	Mult.	Units	Range	Digit	Rate	Rate	Rate

---	--------------	------	------	------	------	------	------	------

00	LED Voltage  1X	2	Volts	0-10	x.xxx	1/10s	1/s	1/m

01	LED Voltage  1Y	2	Volts	0-10	x.xxx	1/10s	1/s	1/m

08	20K Refrig Temp	100	Kelvin	0-360	x.x	1/s	1/s	1/m

09	20K Amp Temp	100	Kelvin	0-360	x.x	1/s	1/s	1/m

0A	70K Refrig Temp	100	Kelvin	0-360	x.x	1/s	1/s	1/m

0B	70K OMT Temp	100	Kelvin	0-360	x.x	1/s	1/s	1/m

0C	FE Box Temp	100	Kelvin	0-360	x.x	1/s	1/s	1/m

0D	Dewar Vacuum 	1	mV	0-9.9	x.x	1/s	1/s	1/m

0E	Vac-Ion Pump Vacuum	Note 1	mV	10E-10	x.xE-x	1/30s	1/s	1/m

to 0

0F	+15 Supply Mon	11.111	Volts	0-20	x.xxx	1/s	1/s	1/m

10	-15 Supply Mon	11.111	Volts	-20-0	x.xxx	1/s	1/s	1/m

11	+5 Supply Mon	1	Volts	0-10	x.xxx	1/s	1/s	1/m

12	-5 Supply Mon	1	Volts	-10-0	x.xxx	1/s	1/s	1/m

13	+28 Supply Mon 4	11.111	Volts	0-30	x.xxx	1/s	1/s	1/m

14	X amp 1, Gate 1V	2	Volts	-20-0	x.xxx	1/10s	1/s	1/m

15	X amp 1, Gate 2V	2	Volts	-20-0	x.xxx	1/10s	1/s	1/m

16	X amp 2, Gate 1V	2	Volts	-20-0	x.xxx	1/10s	1/s	1/m

17	X amp 2, Gate 2V	2	Volts	-20-0	x.xxx	1/10s	1/s	1/m

18	Y amp 1, Gate 1V	2	Volts	-20-0	x.xxx	1/10s	1/s	1/m

19	Y amp 1, Gate 2V	2	Volts	-20-0	x.xxx	1/10s	1/s	1/m

20	Y amp 2, Gate 1V	2	Volts	-20-0	x.xxx	1/10s	1/s	1/m

21	Y amp 2, Gate 2V	2	Volts	-20-0	x.xxx	1/10s	1/s	1/m

24	LO Level	1	Volts	0-10	x.xxx	1/10s	1/s	1/m

28	Y Cal Curr.	10	mA	0-50	x.xx	1/10s	1/s	1/m

29	X Cal Curr.	10	mA	0-50	x.xx	1/10s	1/s	1/m

2A	IF Pwr (x pol)	1	V	0-10	x.xx	10/s	1/s	1/m

2B	IF Pwr (y pol)	1	V	0-10	x.xx	10/s	1/s	1/m

Note1:  Vac-Ion Pump function is

Output =   10 EXP  (0.5313*(Input)-10.250)

Table 11:  Analog Monitor Alarms

Alarms on these functions shall be set as follows:

RA

(hex)	Description	Mult.	Units	Alarm

---	--------------	------	------	------

00	LED Voltage  1X	2	Volts	< 0.5

01	LED Voltage  1Y	2	Volts	< 0.5

08	20K Refrig Temp	100	Kelvin	> 30

09	20K Amp Temp	100	Kelvin	> 30

0A	70K Refrig Temp	100	Kelvin	> 100

0B	70K OMT Temp	100	Kelvin	> 110

0C	FE Box Temp	100	Kelvin	none

0D	Dewar Vacuum 	1000	mV	> 1.0

08	Vac-Ion Pump Vac	Note 1	mV	none

0F	+15 Supply Mon	11.111	Volts	< 13.5, >16

10	-15 Supply Mon	11.111	Volts	> -13.5. < -16

11	+5 Supply Mon	1	Volts	< 4.5, > 5.5

12	-5 Supply Mon	1	Volts	> -4.5, < -5.5

13	+28 Supply Mon 4	11.111	Volts	< 25, > 30

14	X amp 1, Gate 1V	2	Volts	none

15	X amp 1, Gate 2V	2	Volts	none

16	X amp 2, Gate 1V	2	Volts	none

17	X amp 2, Gate 2V	2	Volts	none

18	Y amp 1, Gate 1V	2	Volts	none

19	Y amp 1, Gate 2V	2	Volts	none

20	Y amp 2, Gate 1V	2	Volts	none

21	Y amp 2, Gate 2V	2	Volts	none

24	LO Level	1	Volts	none

28	Y Cal Curr.	10	mA	none

29	X Cal Curr.	10	mA	none

2A	IF Pwr (x pol)	1	V	none

2B	IF Pwr (y pol)	1	V	none

Abbreviations:

A:	Amps

Amp:	Amplifier

Cal:	Calibrate

Curr:	Current

FE:	Front End

K:	Kelvin

LED:	Light Emitting Diode

M:	Minute

MCB:	Monitor and Control Bus

Mon:	Monitor

MSB:	Most Significant Bit

Mult.:	Multiplier

Pol:	Polarization

RA:	Relative Address

Refrig:	Refrigerator

S:	Second

TTL:	Transistor/Transistor Logic

V:	Volts

Vac:	Vacuum