In January, 1995, a time and frequency distribution system was put

into operation.  A "Site 1Hz" tick is generated by a TRAK master clock

from a 5 MHz signal from the Sigma Tau hydrogen maser standard.

The tick and also an IRIG time signal is distributed around the

Green Bank site on fiber optics.  A clock offset measurement system

logs offsets between various clocks about every 10 minutes.  Files

listing these offsets are placed in the anonymous ftp area on osiris

under directory "timer".

Clock offset file names are of the form:  YYYY_MM.tdat, one file per month.

Files are updated daily.  The files contain one entry per hour, which is

a median of values recorded during the hour.  Typically this is a median

of 6 measurements.  If a number such as "987654321" appears, there was no

data recorded for that device.

-----------------------------------------------------

Format of the clock offset tables -  "*.tdat" format:

-----------------------------------------------------

Columns are labeled:   ST-GPS  ST-SAO   GPS-SAO  GPS-F85  GPS-F20

where xxx-yyy means  device xxx minus yyy.

A positive quantity means the pulse from xxx precedes yyy.

ST = Sigma Tau maser (= Site 1Hz)

SAO = SAO maser at 140-ft, corrected for travel time of 12392 ns

from the 140-ft building to the Interferometer building

on the optical fiber.

Note:   ST-SAO is calculated as:  ST-SAO-12392, and

GPS-SAO is calculated as: GPS-SAO-12392.

GPS = DATUM GPS receiver at the Interferometer building.

F85 = The formatter of the 85-ft USNO VLBA recording system.

F20 = The formatter of the 20-meter USNO VLBA recording system.

-----------------------------------------------------

Format of the clock offset tables -  "*.ttac" format:

-----------------------------------------------------

Columns are labeled:   ST-TAC  ST-SAO   TAC-SAO  TAC-F85  TAC-F20

These data are the same as the *.tdat  files, except that the "TAC" GPS receiver

is used instead of the old DATUM.  The TAC is Tom Clark's "Totally Accurate

Clock", based on a Motorola GPS chip.  The TAC tracks 6 satellites at a time,

and gives a more accurate time than the DATUM, which tracks only one.

-----------------------------------------------------

Format of the clock offset tables -  "*.t0" format:

-----------------------------------------------------

Table lists time differences with respect to the Site 1Hz (derived

from the Sigma Tau maser).

Columns are labeled:  Site1Hz, GPS, OldST, 140ft, VLB85, VLB20, 140Rtn

Thus column 1 is time difference: Site1Hz minus Site1Hz

col 2 is Site1Hz minus GPS, and so on.

OldST is 1 HZ derived from the Sigma Tau maser by the old counter

that was used with the HP system.

140ft is 1Hz from the 140ft maser. The cable from the 140-ft to

the Interferometer building has about 5.8 microsec delay.

VLB85 = formatter tick of the 85-ft USNO VLBA data acquisition system.

VLB20 = formatter tick of the 20-meter USNO VLBA data acquisition system.

140Rtn = Site1Hz after being sent to the 140-ft and back on fiber.

This is the round trip delay.

-----------------------------------------------------

Format of the clock offset tables -  "*.t1" format:

-----------------------------------------------------

Table lists time differences with respect to GPS.  Otherwise all is

as described for the "*.t0" tables.